Magnetic Resonance Imaging:

The Art of Sequence Design

E. Mark Haacke, Ph.D.

Department of Radiology, Wayne State University, Detroit, Michigan, USA

Ruijin Hospital, Shanghai Jiao Tong University, Shanghai, China

## Abstract

Designing a magnetic resonance imaging (MRI) pulse sequence is, at its core, an exercise in constrained optimization. A modern gradient echo sequence alone offers at least 25 adjustable parameters — repetition time, echo time, bandwidth, flip angle, spatial resolution, field of view, flow compensation, partial Fourier fraction, parallel imaging factor, and several others — each of which interacts with the others in ways that are rarely intuitive. If each parameter had only two possible settings, the combinatorics alone would generate more than 32 million potential sequence configurations. In practice, physical and clinical constraints — imaging time, spatial resolution, signal-to-noise ratio (SNR), and the contrast mechanism of interest — collapse this enormous space down to a tractable set of a few hundred reasonable choices. This review systematically works through the major parameters that govern sequence design, the physical relationships that link them, and the trade-offs a sequence designer must navigate to satisfy the perennial radiological demand: the fastest possible acquisition, with the best possible SNR, in the shortest possible time.

# 1. Introduction

Every MRI pulse sequence represents a negotiated compromise among competing physical constraints. Although the underlying physics of nuclear magnetic resonance is well established, the practical design of a sequence for a specific clinical or research application requires the simultaneous management of more than twenty interacting variables. These include timing parameters such as repetition time (TR) and echo time (TE); acquisition parameters such as receiver bandwidth (BW) and the number of echoes acquired (Ne); geometric parameters such as spatial resolution (Δx, Δy, Δz) and field of view (FOV); and a range of acceleration and post-processing choices including parallel imaging, partial Fourier sampling, flow compensation, and various filtering strategies.

If each of these variables could independently take on only two values, the resulting design space would exceed 32 million combinations. Fortunately, the space is heavily constrained in practice. Imaging time and spatial resolution requirements immediately rule out large swaths of parameter combinations, and the underlying physics — most notably the Nyquist sampling relationship that links gradient strength, field of view, resolution, and acquisition time — ties many of these variables together so that fixing one fixes several others. The result is that a sequence designer is typically choosing among a few hundred genuinely distinct configurations rather than millions, but doing so wisely still requires a clear, integrated understanding of how every parameter affects image quality, contrast, and acquisition time. Further, what might be considered ideal from a technical perspective may not be absolutely necessary from a clinical perspective, for example, a resolution of 0.25mm might best show boundaries of most structures in the brain but a trade-off of 0.5mm may be sufficient for excellent clinical interpretation of the data. This would save significant time, storage space, and increase signal-to-noise (SNR) at the same time.

The unifying objective, as any practicing radiologist will confirm, is straightforward to state even if difficult to achieve: the fastest possible imaging, with the best possible SNR, in the shortest possible time. Both magnitude and phase images are now routinely used in clinical and research practice — magnitude images for conventional T1-weighted (T1W), T2-weighted, and T2*-weighted (T2*W) contrast, and phase images for quantitative susceptibility mapping (QSM) and susceptibility weighted imaging (SWI). This dual use further constrains and complicates sequence design, since the echo times and other parameters that are optimal for magnitude contrast are not necessarily optimal for phase-based contrast.

This review organizes the major sequence-design parameters into five functional groups: timing parameters that govern contrast and acquisition speed; bandwidth, signal-to-noise, and resolution relationships; spatial encoding and the Nyquist relationship; acceleration and filtering strategies; and contrast-modifying choices such as flip angle, flow compensation, and magnetization transfer. Throughout, the aim is to make explicit the physical dependencies that link these parameters so that a sequence designer can reason about trade-offs systematically rather than by trial and error.

# 2. Timing Parameters: Repetition Time and Echo Time

The repetition time (TR) and echo time (TE) are the two most fundamental timing parameters in any MRI sequence, and together they largely determine the contrast weighting of the resulting image. The main focus in this review is on 3D gradient echo imaging so the use of inversion pulses is not included. These too, of course, and yet further to the complications and flexibility of sequence design.

## 2.1 Repetition Time

A short TR enables rapid three-dimensional imaging, since the total acquisition time scales directly with TR multiplied by the number of phase-encoding steps. However, a short TR also limits how long the echo time can be, since TE must necessarily fall within the TR window. Depending on the flip angle selected, a short-TR acquisition can produce either T1-weighted or proton-density-weighted (PDW) contrast: low flip angles minimize T1 effects and approach pure spin-density weighting, while higher flip angles introduce progressively stronger T1 weighting.

## 2.2 Echo Time

The echo time, or in multi-echo sequences the set of echo times {TEₙ}, is chosen according to the contrast mechanism of interest. T1-weighted imaging requires only a short TE, since the goal is to minimize T2* decay and isolate T1 contrast. By contrast, T2*-weighted imaging, QSM (depending on the susceptibility of the tissue of interest), and SWI all require longer echo times, since the phase accumulation and signal decay that generate these contrasts both increase with TE.

An important practical consequence follows from combining a short TR with a high bandwidth: the resulting images begin to resemble those from a spin echo (SE) acquisition. This occurs for two independent reasons. First, high bandwidth reduces chemical shift and susceptibility-related geometric distortion in the image, bringing the gradient echo geometry closer to the comparatively distortion-free spin echo case. Second, short echo times limit the amount of T2*-related signal loss that can accumulate, which similarly pushes gradient echo contrast toward the spin-echo-like appearance where T2* effects are minimized. (Higher resolution imaging also reduces external field related sources of T2* dephasing such as those from air/tissue interfaces.)

A further timing constraint arises in water/fat imaging, which requires the echo time to be set precisely so that water and fat signals are either in-phase or opposed-phase at the time of readout, since the relative phase between these two species depends on the chemical shift frequency difference and the elapsed time since excitation.

# 3. Bandwidth, Signal-to-Noise Ratio, and the Number of Echoes

## 3.1 Receiver Bandwidth

Receiver bandwidth (BW) governs a fundamental trade-off between SNR and image fidelity. SNR depends on bandwidth as 1/√BW, so a lower bandwidth directly increases SNR. However, lower bandwidth comes at a cost: it lengthens the spacing between echoes in multi-echo sequences, increases geometric distortion (particularly near air/tissue interfaces and metal), and increases the spatial shift of fat signal relative to water in the resulting images, since chemical shift artifact scales inversely with bandwidth. High bandwidth reverses each of these trade-offs — worse SNR in exchange for reduced distortion and reduced fat-shift artifact.

## 3.2 Number of Echoes

In multi-echo acquisitions, the minimum spacing between successive echoes (ΔTE) is set by the receiver bandwidth and by the rise time of the imaging gradients — faster, stronger gradients permit shorter inter-echo spacing. A long ΔTE is important when the goal is to measure T2* accurately, since the precision of a T2* estimate from a multi-echo series improves with the echo spacing, even as the per-echo SNR itself varies with √ΔTE.

# 4. Flip Angle, Spatial Resolution, and Field of View

## 4.1 Flip Angle

The flip angle (FA = θ) is one of the most consequential contrast-determining parameters in gradient echo imaging. A low flip angle produces spin-density-weighted contrast with little to no T1 effect, and has the additional benefit of reducing or eliminating pulsatility artifacts arising from inflowing blood or CSF pulsation. A high flip angle, conversely, produces strong T1-weighted contrast, as the degree of T1 weighting in a spoiled gradient echo sequence increases monotonically with flip angle up to the Ernst angle and beyond.

## 4.2 Spatial Resolution and Voxel Volume

Spatial resolution is described by the voxel dimensions Δx, Δy, and Δz, with voxel volume given by Δv = Δx·Δy·Δz. In typical clinical and research brain imaging, in-plane resolution Δx is on the order of 1 mm, with a practical range extending from approximately 0.1 mm to 1 mm depending on field strength, coil performance, and acquisition time constraints.

## 4.3 Field of View

The field of view in each direction (FOV = L) is determined by the product of the number of samples and the voxel dimension in that direction:

Lx = NxΔx,    Ly = NyΔy,    Lz = NzΔz

These relationships mean that FOV, resolution, and matrix size are not independent choices — fixing any two determines the third, and all three interact with acquisition time through the Nyquist sampling relationship described in the following section.

# 5. The Nyquist Relationship and Spatial Encoding

The Nyquist sampling theorem, applied to the spatial encoding gradients used in MRI, links gradient strength, spatial resolution, field of view, and acquisition time through a small number of governing equations. For the readout (frequency-encoding) direction:

γGxΔxTacq = 2π,    where Tacq = NxΔt

γGxLxΔt = 2π,    where Lx = NxΔx

These equations make explicit that all of the spatial encoding parameters — gradient amplitude, resolution, field of view, sampling interval, and acquisition time — are mutually constrained rather than freely adjustable. A worked numerical example illustrates the practical implications. If the maximum gradient strength is Gx,max = 23.5 mT/m, then γGx ≈ 1 MHz/m, giving Δx·Tacq = 10⁻⁶ m·s = 1 mm·ms. Consequently, if the desired in-plane resolution is Δx = 1 mm, the readout acquisition window must be Tacq = 1 ms. If the field of view is set to Lx = 250 mm, then Nx = 250 samples are required, which in turn fixes the dwell time at Δt = 4 μs. The corresponding bandwidth per pixel is BW/pixel = 1/(NxΔt) = 1000 Hz/pixel, though it should be noted that bandwidth is sometimes quoted instead as BW = 1/Δt across the full field of view Lx, a convention that can cause confusion when comparing bandwidth values across vendors or publications.

Similar expressions for the phase encoding direction along y are given by:

γNΔGyΔyτ = 2π,    where Gymax = NyΔGy/2

γGxLxτ = 2π,    where Ly = NyΔy

where τ is the duration of Gy. Recall that, in the process of phase encoding, a tab file is used to run from -Gymax to Gymax-ΔGy.

Modern gradient hardware substantially expands the achievable design space relative to the example above: contemporary clinical and research gradient systems can reach amplitudes as high as 200 mT/m with slew rates of 200 T/m/s, corresponding to a rise time to peak gradient of approximately 1 ms. These higher-performance gradients permit shorter minimum echo times, shorter inter-echo spacing in multi-echo sequences, higher spatial resolution and the ability to fully flow compensate between echo times, within a given acquisition time than were achievable with earlier-generation gradient hardware.

# 6. The Signal-to-Noise Ratio Equation

SNR ratio in MRI depends jointly on voxel volume, total acquisition time, and receiver bandwidth according to the relationship:

SNR ∝ Δv · √time / √BW

This single relationship governs many of the practical trade-offs a sequence designer must navigate. Increasing the number of signal averages (Nacq) improves SNR only as the square root of the number of averages — doubling acquisition time through averaging (Nacq = 2) yields only a √2-fold SNR gain, an unfavorable trade in scan-time-limited clinical settings. Increasing spatial resolution carries an analogous penalty: if the number of phase-encoding steps in the y-direction (Ny) is doubled to improve resolution in that direction, SNR correspondingly decreases by a factor of √2, all else being equal, since voxel volume has been halved. This is because the voxel volume goes down by a factor of 2 but also the noise is reduced by a factor of √2, therefore, there is only a loss of √2 in SNR.

# 7. Post-Processing Filters: Sliding Window and AI-Based Noise Reduction

## 7.1 Sliding Window and Hanning-like Filters

A sliding window filter (SWF), or related Hanning-type smoothing filter, can be applied after image reconstruction to improve SNR at some cost to effective spatial resolution. A SWF with a 2-2-1 weighting kernel, for example, averages over four pixels and increases SNR by a factor of approximately 2. Such filters are best applied after the three-dimensional dataset has been reformatted into the imaging plane of clinical interest, since the achievable SNR gain depends on the relationship between the acquired voxel dimensions and the reformatted slice thickness. However, there is a major advantage to starting with high resolution and then filtering it, despite the longer scan time, because after applying the SWF 221, the matrix size is invariant making the image appear to have higher resolution and as an added benefit this reduces the Gibbs artifact considerably. If lower resolution data had been take the resulting SNR would be the same but the Gibbs ringing artifact would be clearly present and disruptive in interpreting the data.

A representative example illustrates this dependence. Consider a sagittally acquired CSP (contiguous slice partition) dataset with native voxel dimensions of 0.67 × 1 × 1 mm, interpolated to 0.67 mm isotropic resolution. If this volume is reformatted to the transverse plane, the resulting slice thickness becomes 0.67 mm, and applying a 1-1-3 sliding window filter (SWF113) in this configuration yields an SNR increase of √3. However, if the same dataset is instead viewed in the coronal or sagittal plane — the planes closer to the original acquisition geometry — the same SWF113 filter yields only a √2 SNR gain. This discrepancy arises because the data along the dimension being filtered originated from a native 1 mm acquisition that was subsequently interpolated down to 0.67 mm, meaning the apparent resolution gain in the interpolated dimension does not correspond to genuinely independent samples.

## 7.2 Artificial Intelligence Noise Reduction

Artificial intelligence (AI)-based noise-reducing filters operate on a conceptually similar principle to conventional sliding window and Hanning filters — trading some degree of effective independent spatial information for an improvement in apparent SNR — but achieve this trade-off through learned, data-driven reconstruction or denoising models rather than fixed convolutional kernels. These approaches have become increasingly prevalent in clinical MRI reconstruction pipelines as an alternative or complement to conventional filtering.

# 8. Flow Compensation and Partial Fourier Acquisition

## 8.1 Flow Compensation

Flow compensation (FC) gradients are designed to rephase the signal from moving spins (most commonly flowing blood or pulsatile CSF) that would otherwise dephase and cause signal loss or ghosting artifact. Omitting flow compensation entirely allows for the shortest possible echo time, since FC gradient lobes add extra time to the sequence. Partial flow compensation, applied only in the slice-select and frequency-encoding (read) directions, keeps TE reasonably short while still addressing the most common sources of flow artifact. Full flow compensation, applied simultaneously in the slice, x, and y directions, provides the most complete artifact suppression but typically requires an additional 5–10 ms between echoes to accommodate the necessary gradient lobes.

## 8.2 Partial Fourier Acquisition

Partial Fourier (PF) acquisition exploits the conjugate symmetry of k-space to reduce the number of phase-encoding lines that must be acquired. A PF fraction of 1 corresponds to a fully symmetric echo with no time savings. A PF fraction of 7/8, by contrast, corresponds to an asymmetric echo in which only 3/8 of the data is acquired on one side of the k-space center and 4/8 on the other. When applied in the frequency-encoding (read) direction, partial Fourier sampling allows for a shorter minimum echo time, since less time is required to traverse from the start of the readout to the k-space center. When applied in the phase-encoding directions (y and/or z), partial Fourier sampling instead accelerates the overall acquisition: setting PFy = PFz = 7/8 in both phase-encoding directions saves approximately 25% of total scan time. Using smaller PF fractions is dangerous if the resolution is low but for high resolution imaging it is more forgiving and when applied in phase and partition encoding say with 6/8 in each direction, this can save almost a factor of 2 in time.

# 9. Parallel Imaging and Compressed Sensing Acceleration

Parallel imaging, characterized by an acceleration factor commonly denoted iPAT (integrated Parallel Acquisition Technique), or its compressed sensing (CS) analogue, is used to accelerate data acquisition by exploiting the spatial sensitivity profiles of a multi-channel receiver coil array (in the case of parallel imaging) or the inherent sparsity of the underlying image (in the case of compressed sensing). Acceleration factors of iPAT = 2 or CS = 2 – 2.5 are routine in current clinical practice, and factors of 3 are increasingly common as coil arrays and reconstruction algorithms have matured. For applications such as MR angiography (MRA), where the underlying vascular anatomy is highly sparse, even more aggressive acceleration factors of CS = 6 to 10 are achievable without unacceptable degradation of image quality. Acceleration necessarily comes at a cost to SNR, since fewer raw data samples are acquired per unit time; the practical acceleration factor that can be tolerated in a given application is, therefore, determined by a balance between the desired time savings and the SNR budget available from coil sensitivity, field strength, and voxel size.

# 10. Additional Acquisition Parameters

Beyond the core timing, bandwidth, and acceleration parameters discussed above, several additional choices materially affect both image quality and acquisition time.

## 10.1 Field of View, Number of Acquisitions, and Number of Slices

The choice of field of view, the number of signal averages (Nacq), and the number of slices acquired (which is functionally equivalent to resolution in the slice-select direction for three-dimensional acquisitions) round out the basic geometric and averaging parameters that must be specified for any sequence.

## 10.2 Magnetization Transfer Contrast

A magnetization transfer contrast (MTC) preparation pulse can be added to suppress signal from tissue with a substantial bound macromolecular proton pool. In typical brain imaging applications, an MTC pulse suppresses white matter signal by a factor of approximately 0.6 and gray matter signal by a factor of approximately 0.675, while leaving cerebrospinal fluid (CSF) signal largely unaffected, since CSF lacks the bound macromolecular pool that mediates the magnetization transfer effect. This differential suppression is particularly useful for improving the visualization of neuromelanin-containing structures and for enhancing vessel conspicuity against suppressed background tissue.

## 10.3 Normalization and Bias Correction

Post-processing normalization or bias-field correction filters improve image homogeneity by compensating for the smoothly varying intensity non-uniformity introduced by coil sensitivity profiles and other system-level effects. Beyond their direct effect on image appearance, these corrections make subsequent window and level adjustment substantially easier for the interpreting radiologist, since they remove much of the large-scale intensity gradient that would otherwise compete with genuine tissue contrast for the available display dynamic range.

# 11. Flyback Acquisition and Echo Offset Choices

In multi-echo gradient echo sequences, the designer must choose whether to use a flyback readout scheme. A flyback acquisition samples only the odd-numbered echoes and maintains the same gradient polarity for every readout, which has the useful side effect of providing automatic flow compensation in the readout direction without requiring additional dedicated FC gradient lobes. However, this scheme requires higher-performance gradients if the spacing between odd echoes (ΔTE) is to be kept as short as possible, since the gradient must return (“fly back”) to its starting position between each useful readout rather than simply reversing polarity for the next echo as in a conventional bipolar multi-echo readout.

A related design choice concerns the selection of echo time offsets for water/fat separation techniques, which, as noted in Section 2.2, require specific in-phase or opposed-phase timing relative to the water–fat chemical shift frequency.

# 12. Multi-Flip-Angle Acquisition Strategies

Finally, a sequence can be designed to acquire data at two or more flip angles within a single protocol, an approach exemplified by STAGE (STrategically Acquired Gradient Echo) and related multi-contrast techniques. By combining the spin-density-weighted contrast obtained at a low flip angle with the T1-weighted contrast obtained at a higher flip angle — both acquired from the same underlying k-space trajectory and echo train — these approaches can synthesize multiple clinically useful contrasts, including T1-weighted, proton-density-weighted, and quantitative T1 maps, from a single, time-efficient acquisition, while also enabling the SWI and QSM processing made possible by acquiring images at sufficiently long echo times.

# 13. Conclusion

Pulse sequence design in MRI is fundamentally an optimization problem conducted within a high-dimensional but heavily constrained parameter space. Of the more than twenty variables available to a sequence designer — repetition time, echo time, bandwidth, number of echoes, flip angle, spatial resolution, field of view, flow compensation, partial Fourier fraction, parallel imaging or compressed sensing acceleration, post-acquisition filtering, magnetization transfer preparation, bias correction, flyback readout strategy, and multi-flip-angle acquisition schemes, among others — only a few hundred combinations represent genuinely distinct, clinically sensible protocols once the governing physical relationships and time/resolution/SNR constraints are applied.

The unifying thread connecting all of these choices is the Nyquist sampling relationship, which ties gradient strength, spatial resolution, field of view, and acquisition time together into a small number of governing equations, together with the SNR relationship SNR ∝ Δv·√time/√BW, which quantifies the fundamental trade-off between speed, resolution, and image quality that underlies every sequence design decision. A sequence designer who internalizes these relationships, together with the qualitative effects of flip angle, echo time, and the various acceleration and filtering strategies surveyed here, is well equipped to navigate the genuinely large, but ultimately tractable space of clinically useful MRI sequence designs, and to do so in pursuit of the radiologist's enduring goal: the fastest imaging, with the best SNR, in the shortest time. For a summary of all the key governing equations and variables, see the tables below.

Key governing relationships linking these parameters:

| Relationship | Equation | Implication |
| --- | --- | --- |
| Nyquist (read direction) | γGxΔxTacq = 2π | Fixes Gx, Δx, Tacq together |
| FOV–resolution–matrix | Lx = N_xΔx | Any two determine the third |
| SNR scaling | SNR ∝ Δv · √time / √BW | All resolution, time, and BW trade-offs follow from this |
| Scan time | T_scan = · NyNzNacqTR / (iPAT · PFy · PFz) | Every acceleration parameter enters here |

| # | Parameter | Symbol | Range (3T) | Primary Effect on Image |
| --- | --- | --- | --- | --- |
| 1 | Repeat Time | TR | 20–3000 ms | T1 contrast; scan time; enables short-TE |
| 2 | Echo Time | TE | 2–40 ms | T2* contrast; phase accumulation for SWI/QSM |
| 3 | Number of Echoes | Ne | 1–8 | SNR (averaging); R2* mapping; multi-contrast |
| 4 | Echo Time Spacing | ΔTE | 2–10 ms | R2* fit precision; inter-echo field map quality |
| 5 | Flip Angle | FA (θ) | 6°–30° | T1 weighting; pulsatility artifact; Ernst angle optimization |
| 6 | Receiver Bandwidth | BW | 100–500 Hz/pixel | SNR (∝1/√BW); fat shift; geometric distortion |
| 7 | In-Plane Resolution | Δx, Δy | 0.5–1.0 mm | Resolution; SNR (∝ voxel volume) |
| 8 | Slice Thickness | Δz | 0.67–2.0 mm | Through-plane resolution; background gradient sensitivity |
| 9 | Field of View | FOV (Lx, Ly, Lz) | 200–260 mm | Coverage; aliasing; determines N given Δx |
| 10 | Matrix Size | Nx, Ny, Nz | 256–512 | Resolution × FOV product; scan time |
| 11 | Number Slices | Nslices | 40–120 | Volume coverage; scan time |
| 12 | Number Averages | Nacq | 1–4 | SNR (∝√N_acq); scan time penalty |
| 13 | Partial Fourier Fraction (Read) | PFread | 6/8–1 | Shortens minimum TE; mild SNR reduction |
| 14 | PF phase/ partition | PFy, PF_z | 6/8–1 | Reduces scan time ~25% at 7/8; mild SNR/resolution trade |

| # | Parameter | Symbol | Range (3T) | Primary Effect on Image |
| --- | --- | --- | --- | --- |
| 15 | Parallel Imaging / Compressed Sensing Factor | iPAT / CS | 1–3 (routine); up to 6–10 (MRA) | Accelerates acquisition; SNR penalty ∝ √factor |
| 16 | Flow Comp gradients | FC | None / Partial (read+slice) / Full (x,y,z) | Reduces pulsatility artifact; adds 5–10 ms to TE per direction |
| 17 | Flyback Readout | FB | On / Off | Auto flow-comp in read direction; requires higher gradient performance |
| 18 | Sliding Window / Hanning Filter | SWF | Off / SWF113 / SWF221 | SNR gain (√2 to √3); slight resolution reduction |
| 19 | AI Noise Reduction Filter | NRF | Off / On | SNR improvement; data-driven; trades independent spatial information |
| 20 | Magnetization Transfer Contrast Pulse | MTC | On / Off | WM suppressed ×0.6; GM ×0.675; CSF unaffected; improves vessel conspicuity |
| 21 | Bias / Normalization Filter | BIAScorr | On / Off | Corrects coil sensitivity non-uniformity; improves window/level display |
| 22 | Multi-Flip-Angle | FA₁, FA₂ | e.g. 6° + 30° | Enables multi-contrast (PDW + T1W + SWI/QSM) from STAGE |
| 23 | Water/Fat Echo Offset | TE offset | In-phase / Opposed-phase | Required for water-fat imaging; constrains TE choices |
| 24 | Gradient Strength / Slew Rate | G_max, SR | Up to 200 mT/m; 200 T/m/s | Minimum TE; minimum ΔTE; spatial resolution at fixed acquisition time |
| 25 | Addition of pre-pulses | N/A | N/A | Ability to modify contrast as in FLAIR and MP-RAGE |